Name of Equipment:  	Property #:

Version or Model:	Clinical Engineering #:

Serial #:	Lab ID #:

Function of Equipment:	Department of use:

Date of Receipt	Purchase Price:

Implementation Time Line:

|  | Date completed / NA | Comments | Initials |
| --- | --- | --- | --- |
| Installation: |  |  |  |
| Unpack and inspect equipment | 09-09-2016 | Vendor Lab |  |
| Ensure that equipment is installed per manufacturer specifications | 09-09-2016 | Upright and out of box, wait 24 hours prior to turning on. At least 4 inches clearance for ventilation |  |
| Contact Clinical Engineering or Engineering Services, as applicable, to perform electrical/safety checks |  |  |  |
| Ensure seismic compliance is met | NA | Equipment is portable and on a cart |  |
| Attach asset tag (Organization property number) to equipment, and record number on top of this form, as well as on Memorandum from Organization Finance (give copy of Memorandum to Administration) |  | Email to Finance 09/12/16 requesting asset tags |  |
| Assign Organization Clinical Laboratory ID number and record on the equipment, on the top of this form, and update the Clinical Laboratory Equipment Master List (S/Midman/Equipment/Equipment Inventory) | 09/12/16 RF-106 |  |  |
| SOPs and Forms: |  |  |  |
| Draft or update applicable SOPs |  |  |  |
| Draft or update applicable training/competency documents |  |  |  |
| Determine QC and preventive maintenance schedule and create or update forms, as applicable |  |  |  |
| Assess the need for inter-instrument or inter-method comparison on an ongoing basis. |  |  |  |
| Update departmental task lists to include QC and/or preventive maintenance of new equipment, as applicable |  |  |  |

| Update inventory lists with equipment specific reagents/supplies, as applicable | NA |  |  |
| --- | --- | --- | --- |

| # Operation and Performance: list equipment specific validation tasks to be performed according to manufacturer instructions | Expected Results | Observed Results | Acccept-able? | Date | Ini. |
| --- | --- | --- | --- | --- | --- |
| Turning on Refrigerator Turn Key to ON position | Refrigerator turns on, operating on internal DC battery | Temperature display comes on | Y N |  |  |
| Plug AC cord from refrigerator into wall outlet | AC operation is indicated by the green/red LED on the battery charger. Digital temperature display will be on. |  | Y N |  |  |
| Programming the Medical Refrigerator: Press and hold the PRG button until P1 appears, then release Use Arrow Keys to scroll, and ensure that P1, P25, P26, P27, P28 are present in the programming menu | Arrow keys appear, allowing access to various parameters P1, P25, P26, P27, P28 are present |  | Y N Y N |  |  |
| Scroll to P1, press “Set”, use arrow keys to adjust the value to 1, press “Set” to confirm Scroll to P25, press “Set” use arrow keys to adjust the value to 1.5, press “Set” to confirm Scroll to P26, press “Set” use arrow keys to adjust the value to 2, press “Set” to confirm Scroll to P27, press “Set” use arrow keys to adjust the value to 8, press “Set” to confirm Scroll to P28, press “Set” use arrow keys to adjust the value to 0, press “Set” to confirm | Able to set all parameters to the values indicated in Section 4, Table 1, Instruction Manual P28 can be set in higher minute value to delay alarm response |  | Y N |  |  |
| Set the Thermostat Temperature: Press and hold “Set” button until display shows ST1 Release “Set” button and current thermostat set point will flash Press either Up or Down Arrow to change the set point to 3.0C Press “Set” to confirm. | ST1 displays Current set point flashes Able to set the Set Point Temperature to 3.0C Display returns to current interior temperature |  | Y N |  |  |
| Allow the unit to run at least one hour or until a stable temperature between 1-6C is reflected on the temperature display. | Within 1 hour, temperature is between 2-8C |  | Y N |  |  |
| Place a verified thermometer inside the unit and allow it to equilibrate at least 10-15 minutes | The verified thermometer should match the temperature display within +1C | Display= Thermometer = | Y N |  |  |
| Perform High and Low Alarm Checks: High: Access Parameter P28 and change the Alarm time allowance to 0 minutes. Access the Set Point and change the set point to 10C. When alarm sounds, note the temperature on the display and on the verified thermometer inside the refrigerator Low: Access Parameter P28 and ensure Alarm time allowance is 0 minutes. Access the Set Point and change the set point to -5C. When alarm sounds, note the temperature on the display and on the verified thermometer inside the refrigerator Verify that Alarms are silenced by pressing the PRG/Mute button Press PRG and UP Arrow together for >5 seconds to reset the alarm. | High Alarm sounds at 8C Low Alarm sounds at 2C Alarm is silenced when PRG/mute is pressed Message ‘rES’ displays to indicate the alarms have been reset, alarm delays are reactivated | When High Alarm Sounds: Display Temp = Thermometer = When Low Alarm Sounds: Display Temp = Thermometer = | Y N Y N Y N Y N |  |  |
| Verify Alarm Delays Change Alarm Delays to 2 minutes for regular use: Press PRG until P1 appears, then release Use up arrow to scroll to P28, press Set, use up or down arrows to set Alarm Delay to 2 minutes Change the Set Point to 10C and watch the temperature display. Start a timer when the display reaches the 8C Change the Set Point to -5C and watch the temperature display. Start a timer when the display reaches 2C | Able to set alarm delay to 2 min Alarm sounds after 2 minutes at 8C Alarm sounds after 2 minutes at 2C | Time at which high alarm sounds = Time at which low alarm sounds = | Y N Y N |  |  |
| Verify Battery Back-up is working: Unplug the unit and check the light on the battery charger and the digital display Leave the unit unplugged for 2 hours Verify that the Battery Back-up indicator light turns red when battery is low (below 12V), and a that a 4 beep high pitched tone sounds Plug into an outlet and verify that the Battery begins charging | The green/red LED on the battery charger will turn off. Digital temperature display will remain on. Back-Up Battery lasts aproximately 48 hours Indicator light turns red and 4 beep tone is heard Battery begins charging when again plugged into an outlet |  | Y N Y N Y N Y N |  |  |

Equipment Validation Plan Defined by:

Technical Supervisor:_________________________________Date_____________

Equipment Validation Plan Approved by:

Chief of Section_______________________________________Date____________

____________________________________________________________________________________________________________________

Following Validation:  Technical Supervisor Review and Approval

All Results Were Acceptable		Date/Initial________________

Some Results Were Unacceptable		Date/Initial________________

Corrective Actions:

Resolution of Unexpected Results

Conclusions and Limitations:

Date/Initial________________

Attach copies of Draft SOP and training documents, if applicable.  Submit for signature and approval.

Equipment validation results are acceptable.  Equipment meets predetermined specifications and quality attributes and performs as expected for its intended use in a normal working environment.

Technical Supervisor___________________________________________________Date_____________

Director, Medical Laboratory _____________________________________________Date ____________

Coordinator-Laboratory, Business & Compliance _____________________________Date_____________

Executive Director______________________________________________________Date_____________

Chief of Section       ____________________________________________________Date_____________

Comments: